BAD: ADCP’s data acquisition is not in stable period, e.g. not per 4 seconds once (6 sec, 8 sec …), or occasionally, the highlights in the GUI window turn red.

When ADCP in slave mode:

1. Trigger out signal must be a clean square wave; there is no strong

noise or ripple (more than 1 volt) on it. Otherwise false trigger

signals will cause the ADCP to be ping incorrectly.

Advice: Put a low pass filter on trigger out cable.

2. Trigger out signal should be at TTL level (0V, 5V). If trigger

signal level is lower than TTL level. The ADCP could skip some

trigger signals.

Advice: A half-bridge driver can be used to strong then the trigger signal to ADCP. Inverters may not work for EK60’s trigger out, if their input impedance is not high enough.

3. ADCP’s trigger in BNC connector must ground well.